## Carrier1-500-CO2
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Calculated | Gas Factor | Calculated |  | voltage span fraction | Air:CO2 factor
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Carrier |  | Carrier |  | (%) | from cubic fit to Hi-Tec calcs |  | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Flowrate Using FFID |  | Flowrate Using FFID |  |  |  |  | fit | fit |  | cubic | linear
Air
2015-03-16 | 14:26 | 3509360797 | 49.6541 | 0.27842 | 1014.945 | 18.9162 | 470.81 | 50.619 | 0.73564 | 37.2376 | 0.74994 | 5.5684 | 1.3594 |  | 37.4113 | 39.6736 |  | 0.46639 | 6.5419
2015-03-16 | 14:28 | 3509360893 | 99.4031 | 0.5497 | 1014.945 | 18.9187 | 1451.547 | 101.1001 | 0.73519 | 74.3277 | 0.74774 | 10.994 | 1.3602 |  | 74.0241 | 74.057 |  | -0.40837 | -0.36413
2015-03-16 | 14:29 | 3509360984 | 199.6915 | 1.1117 | 1014.945 | 18.9212 | 653.6 | 200.5779 | 0.73442 | 147.3085 | 0.73768 | 22.234 | 1.3616 |  | 147.4073 | 145.2877 |  | 0.067082 | -1.3718
2015-03-16 | 14:31 | 3509361070 | 299.7041 | 1.6887 | 1014.945 | 18.9515 | 1199.538 | 299.8417 | 0.73397 | 220.075 | 0.73431 | 33.774 | 1.3625 |  | 220.24 | 218.4196 |  | 0.074982 | -0.7522
2015-03-16 | 14:33 | 3509361193 | 399.817 | 2.2766 | 1014.945 | 18.9771 | 1272.159 | 399.52 | 0.73399 | 293.2424 | 0.73344 | 45.532 | 1.3624 |  | 293.049 | 292.933 |  | -0.065926 | -0.1055
2015-03-16 | 14:34 | 3509361261 | 499.696 | 2.8662 | 1014.945 | 18.9586 | 1748.383 | 498.009 | 0.73461 | 365.8418 | 0.73213 | 57.324 | 1.3613 |  | 365.9011 | 367.6619 |  | 0.016219 | 0.49751
[13 empty] | cubic | linear |  | Hi-Tec calculator output for F-201AC |  |  | cubic | linear
[16 empty] | x | y
[12 empty] | Hi-tec | Carrier 1 |  |  | 10 | 1.36 |  | 1.3584 | 1.3657 | A
[12 empty] | Calibration | High |  |  | 25 | 1.362 |  | 0.00017346 | -0.00016792 | Bx
[12 empty] | names | CO2 |  |  | 50 | 1.362 |  | -7.9187e-07 |  | Cx^2
[12 empty] | A | -1.1461 | 4.3853 |  | 75 | 1.357 |  | -2.3823e-08 |  | Dx^3
[12 empty] | Bx | 140.4277 | 126.745 |  | 100 | 1.344
[12 empty] | Cx^2 | -7.2593
[12 empty] | Dx^3 | 1.0273
[12 empty] | Calib Temp Deg C | 18.9405
[12 empty] | Calib Press Mbars | 1014.945
[12 empty] | Date | 2015-03-16

## Carrier1-500-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU |  |  | Calculated
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source |  |  | Carrier | [5 empty] | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure |  |  | Flowrate Using FFID | [5 empty] | fit | fit |  | cubic | linear
2015-03-16 | 14:26 | 3509360797 | 49.6541 | 0.27842 | 1014.945 | 18.9162 | 470.81 |  |  | 50.619 | 1.0194 |  |  |  |  | 50.8598 | 53.9784 |  | 0.47563 | 6.6367
2015-03-16 | 14:28 | 3509360893 | 99.4031 | 0.5497 | 1014.945 | 18.9187 | 1451.547 |  |  | 101.1001 | 1.0171 |  |  |  |  | 100.6798 | 100.809 |  | -0.41572 | -0.28797
2015-03-16 | 14:29 | 3509360984 | 199.6915 | 1.1117 | 1014.945 | 18.9212 | 653.6 |  |  | 200.5779 | 1.0044 |  |  |  |  | 200.713 | 197.8259 |  | 0.067365 | -1.372
2015-03-16 | 14:31 | 3509361070 | 299.7041 | 1.6887 | 1014.945 | 18.9515 | 1199.538 |  |  | 299.8417 | 1.0005 |  |  |  |  | 300.0732 | 297.4323 |  | 0.077203 | -0.80356
2015-03-16 | 14:33 | 3509361193 | 399.817 | 2.2766 | 1014.945 | 18.9771 | 1272.159 |  |  | 399.52 | 0.99926 |  |  |  |  | 399.2502 | 398.9203 |  | -0.067523 | -0.1501
2015-03-16 | 14:34 | 3509361261 | 499.696 | 2.8662 | 1014.945 | 18.9586 | 1748.383 |  |  | 498.009 | 0.99662 |  |  |  |  | 498.0917 | 500.7018 |  | 0.016609 | 0.54071
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | High
[12 empty] | names | Air
[12 empty] | A | -1.5146 | 5.9153
[12 empty] | Bx | 190.5584 | 172.628
[12 empty] | Cx^2 | -9.1189
[12 empty] | Dx^3 | 1.2036
[12 empty] | Calib Temp Deg C | 18.9405
[12 empty] | Calib Press Mbars | 1014.945
[12 empty] | Date | 2015-03-16
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Qtrace | FF (C) | FF (C) | GMU | GMU | Calculated | GMU | GMU | FF (C) | Cppm | Hi-Tec/FID | Q | c'/C | FF (C) | C | Multi Run | File | GMU | GMU | FF (C) | No Of | Flow | Inst | GMU | Clean_Air | Calculated
of | Measurement | since | Trace | Trace | NPL | Therm | Source | From | Mean | Variance | Trace | Carrier 1 | Carrier | Source | Source | Calibration | Hi-Tec | % | Source |  | Standard Error | % | Start Time |  | Carrier 1 | Carrier 2 | Calibration | Measurement | Measurement |  | Source | Hi-Tec | Carrier
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | FID | C | c^2 | Dem Volt | Meas Q | Flowrate | Dem Strength | Meas Strength | Drift |  |  | l/min |  | in Mean | Error |  |  | Meas Volt | Meas Volt | Filename | Blocks | Time |  | Meas Q | Meas Volt | Flowrate
2015-03-16 | 14:26 | 3509360797 | 0.024992 | 2.2565 | 1014.945 | 18.9162 | 470.81 | 0.024516 | 493.728 | 36.9886 | 2.2573 | 49.6541 | 50.619 | 0.05 | 0.05031 | 0.12 | 503.1 | 1.8982 | 49.6791 | 0.012318 | 1.5112 | 0.30608 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 0.27842 | 0.011009 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 49.6791 | 1.9966 | 50.619
2015-03-16 | 14:28 | 3509360893 | 0.05003 | 4.506 | 1014.945 | 18.9187 | 1451.547 | 0.04919 | 494.856 | 33.8346 | 4.5034 | 99.4031 | 101.1001 | 0.05 | 0.05031 | 0.12 | 503.1 | 1.6659 | 99.4531 | 0.011754 | 1.6775 | 0.339 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 0.5497 | 0.007 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 99.4531 | 1.9978 | 101.1001
2015-03-16 | 14:29 | 3509360984 | 0.05003 | 4.5061 | 1014.945 | 18.9212 | 653.6 | 0.04981 | 249.4293 | 16.3581 | 4.5034 | 199.6915 | 200.5779 | 0.025 | 0.025048 | 0.12 | 250.48 | 0.4212 | 199.7415 | 0.016215 | 0.837 | 0.3356 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 1.1117 | 0.005474 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 199.7415 | 1.9982 | 200.5779
2015-03-16 | 14:31 | 3509361070 | 0.04502 | 4.071 | 1014.945 | 18.9515 | 1199.538 | 0.045 | 150.1459 | 2.5672 | 4.0697 | 299.7041 | 299.8417 | 0.015 | 0.015018 | 0.12 | 150.18 | 0.022711 | 299.7491 | 0.010671 | 0.3389 | 0.22571 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 1.6887 | 0.004721 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 299.7491 | 1.998 | 299.8417
2015-03-16 | 14:33 | 3509361193 | 0.05202 | 4.675 | 1014.945 | 18.9771 | 1272.159 | 0.05206 | 130.2062 | 3.7283 | 4.6736 | 399.817 | 399.52 | 0.013 | 0.013009 | 0.12 | 130.09 | -0.08924 | 399.869 | 0.014829 | 0.3714 | 0.28524 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 2.2766 | 0.004 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 399.869 | 1.9978 | 399.52
2015-03-16 | 14:34 | 3509361261 | 0.05004 | 4.5061 | 1014.945 | 18.9586 | 1748.383 | 0.05021 | 100.4801 | 2.3572 | 4.5032 | 499.696 | 498.009 | 0.01 | 0.010012 | 0.12 | 100.12 | -0.3584 | 499.746 | 0.01528 | 0.3256 | 0.324 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 2.8662 | 0.00384 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 499.746 | 1.9976 | 498.009

## Trace 300
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Trace | Trace | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-01-26 | 12:43 | 3505120982 | 0.25003 | 4.1576 | 1015.839 | 17.835 | 344.341 |  |  | 0.2516 | 1.0063 | bubble flow 1-250cc |  |  | 0.25155 | 0.25125 | -0.019723 | -0.14029
2015-01-26 | 12:45 | 3505121147 | 0.19988 | 3.3132 | 1015.839 | 17.7434 | 344.243 |  |  | 0.1996 | 0.9986 | bubble flow 1-250cc |  |  | 0.19976 | 0.2006 | 0.078841 | 0.50188
2015-01-26 | 12:48 | 3505121309 | 0.1501 | 2.4515 | 1015.839 | 17.8484 | 344.175 |  |  | 0.149 | 0.99267 | bubble flow 1-250cc |  |  | 0.14882 | 0.14892 | -0.11813 | -0.054468
2015-01-26 | 12:50 | 3505121430 | 0.09995 | 1.5898 | 1015.996 | 17.732 | 344.139 |  |  | 0.09817 | 0.98219 | bubble flow 1-250cc |  |  | 0.098187 | 0.097236 | 0.017631 | -0.95147
2015-01-26 | 12:52 | 3505121545 | 0.05008 | 0.7591 | 1016.137 | 17.8065 | 344.218 |  |  | 0.048 | 0.95847 | bubble flow 1-250cc |  |  | 0.048106 | 0.047412 | 0.21994 | -1.2243
2015-01-26 | 12:54 | 3505121690 | 0.024987 | 0.3566 | 1016.137 | 17.8 | 344.238 |  |  | 0.02295 | 0.91848 | bubble flow 1-250cc |  |  | 0.022944 | 0.023271 | -0.026123 | 1.3999
2015-01-26 | 12:57 | 3505121867 | 0.018716 | 0.25909 | 1016.137 | 17.8756 | 344.186 |  |  | 0.01679 | 0.89709 | bubble flow 1-250cc |  |  | 0.016732 | 0.017423 | -0.3429 | 3.7691
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Trace
[12 empty] | Calibration | 300 ml/min
[12 empty] | names | Propane
[12 empty] | A | -2.2e-05 | 0.0018832
[12 empty] | Bx | 0.065407 | 0.059978
[12 empty] | Cx^2 | -0.00297
[12 empty] | Dx^3 | 0.000431
[12 empty] | Calib Temp Deg C | 17.8058
[12 empty] | Calib Press Mbars | 1015.9891
[12 empty] | Date | 2015-01-26

## Trace 50
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Factors |  |  |  |  | Start | End
of | Measurement | since | Trace | Trace | NPL | Therm | Source | FID | Bubble | Flowrate |  |  | Gas | Gas |  | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Calib |  |  |  |  | Reading | Reading |  | fit | fit |  | cubic | linear
2015-02-03
2015-01-23 | 8:50 | 3504847847 | 0.05498 | 4.6458 | 1019.017 | 18.152 | 384.72 | 1 | 0.05134 | 0.05134 | 0.93379 | bubble flow 1-250cc |  |  |  | 0.051333 | 0.051174 |  | -0.013779 | -0.32245
2015-01-23 | 8:53 | 3504847992 | 0.03997 | 3.4128 | 1019.017 | 17.9896 | 384.353 | 1 | 0.03735 | 0.03735 | 0.93445 | bubble flow 1-250cc |  |  |  | 0.037384 | 0.037596 |  | 0.091993 | 0.65777
2015-01-23 | 8:55 | 3504848118 | 0.029987 | 2.558 | 1019.017 | 17.9799 | 384.12 | 1 | 0.02815 | 0.02815 | 0.93874 | bubble flow 1-250cc |  |  |  | 0.028083 | 0.028182 |  | -0.23673 | 0.11345
2015-01-23 | 8:57 | 3504848246 | 0.019986 | 1.6913 | 1019.017 | 17.9845 | 383.95 | 1 | 0.01867 | 0.01867 | 0.93415 | bubble flow 1-250cc |  |  |  | 0.018718 | 0.018637 |  | 0.25916 | -0.17599
2015-01-23 | 8:59 | 3504848374 | 0.009981 | 0.8298 | 1019.017 | 18.1374 | 383.768 | 1.015 | 0.00913 | 0.0092669 | 0.92846 | bubble flow 1-250cc | Corrrected reading |  |  | 0.0092629 | 0.0091496 |  | -0.043667 | -1.2661
2015-01-23 | 9:01 | 3504848501 | 0.004985 | 0.407 | 1019.017 | 18.1959 | 383.581 | 1.045 | 0.004341 | 0.0045363 | 0.91 | bubble flow 1-250cc | Corrrected reading |  |  | 0.0045045 | 0.0044934 |  | -0.70173 | -0.94657
2015-01-23 | 9:03 | 3504848629 | 0.0029761 | 0.23886 | 1019.017 | 18.2026 | 383.353 | 1.075 | 0.002382 | 0.0025606 | 0.8604 | bubble flow 1-250cc | Corrrected reading |  |  | 0.0025827 | 0.0026417 |  | 0.85923 | 3.1657
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Trace
[12 empty] | Calibration | 50 ml/min
[12 empty] | names | Propane
[12 empty] | A | -0.00018089 | 1.1196e-05
[12 empty] | Bx | 0.011657 | 0.011013
[12 empty] | Cx^2 | -0.0003786
[12 empty] | Dx^3 | 5.5142e-05
[12 empty] | Calib Temp Deg C | 18.0917
[12 empty] | Calib Press Mbars | 1019.017
[12 empty] | Date | 2015-02-03 |  | See correction below
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Qtrace | FF (C) | FF (C) | GMU | GMU | GMU | GMU | GMU | FF (C) | Cppm | Hi-Tec/FID | Q | c'/C | FF (C) | C | Multi Run | File | GMU | GMU | FF (C) | No Of | Flow | Inst | GMU | Clean_Air
of | Measurement | since | Trace | Trace | NPL | Therm | Source | From | Mean | Variance | Trace | Carrier 1 | Carrier 2 | Source | Source | Calibration | Hi-Tec | % | Source |  | Standard Error | % | Start Time |  | Carrier 1 | Carrier 2 | Calibration | Measurement | Measurement |  | Source | Hi-Tec
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | FID | C | c^2 | Dem Volt | Meas Q | Meas Q | Dem Strength | Meas Strength | Drift |  |  | l/min |  | in Mean | Error |  |  | Meas Volt | Meas Volt | Filename | Blocks | Time |  | Meas Q | Meas Volt
2015-02-03 | 14:51 | 3505819885 | 0.0024954 | 0.24612 | 1001.639 | 16.9858 | 305.7969 | 0.002701 | 540.23 | 25.4306 | 0.24649 | 5 | 0 | 0.05 | 0.04989 | 0.57 | 498.9 | -7.6504 | 5.0022 | 0.009335 | 0.12133 | 0.022459 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 4.6065 | 0.014186 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 5.0022 | 2.1993
2015-02-03 | 14:51 | 3505819885 | 0.0024954 | 0.24612 | 1001.639 | 16.9858 | 305.7969 | 0.0027042 | 540.865 | 15.6727 | 0.24649 | 5 | 0 | 0.05 | 0.04989 | 0.55 | 498.9 | -7.7589 | 5.0022 | 0.00732 | 0.8838 | 0.1634 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 4.6065 | 0.014186 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 5.0022 | 2.1993
2015-02-03 | 14:47 | 3505819653 | 0.002496 | 0.24612 | 1001.552 | 17.0322 | 304.5289 | 0 | 539.392 | 36.9649 | 0.24641 | 0 | 5.0176 | 0.05 | 0.04972 | 0.57 | 497.2 | -7.8221 | 5.0201 | 0.011272 | 1.0104 | 0.18732 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 0.4639 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 5.0201 | 2.1996
2015-02-03 | 14:47 | 3505819653 | 0.002496 | 0.24612 | 1001.552 | 17.0322 | 304.5289 | 0 | 540.217 | 26.9679 | 0.24641 | 0 | 5.0176 | 0.05 | 0.04972 | 0.55 | 497.2 | -7.9629 | 5.0201 | 0.009613 | 1.3015 | 0.24092 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 0.4639 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 5.0201 | 2.1996
2015-02-03 | 14:30 | 3505818641 | 0.0024961 | 0.24612 | 1001.44 | 17.02 | 299.2919 | 0 | 1074.615 | 57.8929 | 0.24643 | 0 | 2.5068 | 0.1 | 0.09947 | 0.57 | 994.7 | -7.4366 | 2.5093 | 0.00708 | 0.09824 | 0.009142 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.007759 | 0.23934 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 2.5093 | 2.1893
2015-02-03 | 14:30 | 3505818641 | 0.0024961 | 0.24612 | 1001.44 | 17.02 | 299.2919 | 0 | 1077.925 | 13.1564 | 0.24643 | 0 | 2.5068 | 0.1 | 0.09947 | 0.55 | 994.7 | -7.7209 | 2.5093 | 0.003365 | 0.7455 | 0.06916 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.007759 | 0.23934 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 2.5093 | 2.1893
2015-02-03 | 14:59 | 3505820379 | 0.002497 | 0.24617 | 1001.639 | 17.1149 | 309.2647 | 0 | 267.9811 | 21.9387 | 0.24635 | 0 | 9.9942 | 0.025 | 0.024978 | 0.57 | 249.78 | -6.7919 | 9.9967 | 0.017478 | 0.8536 | 0.3185 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.009 | 0.9206 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 9.9967 | 2.1994
2015-02-03 | 14:59 | 3505820379 | 0.002497 | 0.24617 | 1001.639 | 17.1149 | 309.2647 | 0 | 269.2991 | 19.1988 | 0.24635 | 0 | 9.9942 | 0.025 | 0.024978 | 0.55 | 249.78 | -7.2481 | 9.9967 | 0.016271 | 0.8904 | 0.3306 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.009 | 0.9206 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 9.9967 | 2.1994
2015-02-03 | 14:36 | 3505818984 | 0.0029896 | 0.28794 | 1001.441 | 16.9541 | 301.9158 | 0.00319 | 1063.67 | 96.9885 | 0.28882 | 3 | 0 | 0.1 | 0.09959 | 0.57 | 995.9 | -6.3713 | 3.0018 | 0.009259 | 0.14468 | 0.013602 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 2.7956 | 0.013942 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 3.0018 | 2.1992
2015-02-03 | 14:36 | 3505818984 | 0.0029896 | 0.28794 | 1001.441 | 16.9541 | 301.9158 | 0.003198 | 1066.334 | 53.7297 | 0.28882 | 3 | 0 | 0.1 | 0.09959 | 0.55 | 995.9 | -6.6052 | 3.0018 | 0.006874 | 2.0121 | 0.18869 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 2.7956 | 0.013942 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 3.0018 | 2.1992
2015-02-03 | 14:33 | 3505818813 | 0.00499 | 0.4584 | 1001.44 | 16.9737 | 305.8354 | 0.005177 | 1035.868 | 70.0218 | 0.4592 | 5 | 0 | 0.1 | 0.09975 | 0.57 | 997.5 | -3.7039 | 5.0024 | 0.008078 | 0.10804 | 0.01043 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 4.6039 | 0.014954 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 5.0024 | 2.1993
2015-02-03 | 14:33 | 3505818813 | 0.00499 | 0.4584 | 1001.44 | 16.9737 | 305.8354 | 0.005191 | 1038.621 | 31.064 | 0.4592 | 5 | 0 | 0.1 | 0.09975 | 0.55 | 997.5 | -3.9592 | 5.0024 | 0.005366 | 1.0529 | 0.10137 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 4.6039 | 0.014954 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 5.0024 | 2.1993
2015-02-03 | 14:44 | 3505819480 | 0.006244 | 0.5661 | 1001.44 | 17.0814 | 311.0617 | 0 | 512.93 | 32.5557 | 0.5664 | 0 | 12.4938 | 0.05 | 0.04995 | 0.57 | 499.5 | -2.6183 | 12.5 | 0.011124 | 0.6617 | 0.129 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.009 | 1.1552 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 12.5 | 2.1993
2015-02-03 | 14:44 | 3505819480 | 0.006244 | 0.5661 | 1001.44 | 17.0814 | 311.0617 | 0 | 514.279 | 23.2772 | 0.5664 | 0 | 12.4938 | 0.05 | 0.04995 | 0.55 | 499.5 | -2.8737 | 12.5 | 0.009381 | 0.7558 | 0.14696 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.009 | 1.1552 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 12.5 | 2.1993
2015-02-03 | 14:56 | 3505820213 | 0.006246 | 0.5661 | 1001.639 | 17.2084 | 321.188 | 0 | 252.3462 | 15.0882 | 0.5664 | 0 | 24.9999 | 0.025 | 0.02498 | 0.57 | 249.8 | -1.009 | 25.0062 | 0.015393 | 0.5232 | 0.20733 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.01 | 2.359 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 25.0062 | 2.1996
2015-02-03 | 14:56 | 3505820213 | 0.006246 | 0.5661 | 1001.639 | 17.2084 | 321.188 | 0 | 253.731 | 12.0824 | 0.5664 | 0 | 24.9999 | 0.025 | 0.02498 | 0.55 | 249.8 | -1.5493 | 25.0062 | 0.013699 | 0.5312 | 0.20936 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.01 | 2.359 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 25.0062 | 2.1996
2015-02-03 | 14:27 | 3505818463 | 0.006301 | 0.571 | 1001.44 | 17.0613 | 306.2496 | 0 | 1031.038 | 73.1765 | 0.5708 | 0 | 6.2871 | 0.1 | 0.10012 | 0.57 | 1001.2 | -2.894 | 6.2934 | 0.008297 | 0.12382 | 0.012009 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.007754 | 0.5789 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 6.2934 | 2.1995
2015-02-03 | 14:27 | 3505818463 | 0.006301 | 0.571 | 1001.44 | 17.0613 | 306.2496 | 0 | 1034.106 | 32.2055 | 0.5708 | 0 | 6.2871 | 0.1 | 0.10012 | 0.55 | 1001.2 | -3.1821 | 6.2934 | 0.005488 | 1.0932 | 0.10571 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.007754 | 0.5789 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 6.2934 | 2.1995
2015-02-03 | 14:24 | 3505818295 | 0.012498 | 1.113 | 1001.441 | 17.1233 | 311.0075 | 0 | 1010.503 | 262.161 | 1.1128 | 0 | 12.4931 | 0.1 | 0.09994 | 0.57 | 999.4 | -1.0988 | 12.5056 | 0.016023 | 2.5444 | 0.2518 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.007485 | 1.1549 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 12.5056 | 2.1995
2015-02-03 | 14:24 | 3505818295 | 0.012498 | 1.113 | 1001.441 | 17.1233 | 311.0075 | 0 | 1012.684 | 208.98 | 1.1128 | 0 | 12.4931 | 0.1 | 0.09994 | 0.55 | 999.4 | -1.3118 | 12.5056 | 0.014275 | 2.5919 | 0.25594 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.007485 | 1.1549 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 12.5056 | 2.1995
2015-02-03 | 14:54 | 3505820048 | 0.012498 | 1.113 | 1001.639 | 17.2067 | 350.919 | 0 | 249.431 | 3.2745 | 1.114 | 0 | 50.0405 | 0.025 | 0.02497 | 0.57 | 249.7 | 0.10785 | 50.053 | 0.007255 | 0.026537 | 0.010639 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.015482 | 4.67 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 50.053 | 2.1994
2015-02-03 | 14:54 | 3505820048 | 0.012498 | 1.113 | 1001.639 | 17.2067 | 350.919 | 0 | 250.7893 | 1.2783 | 1.114 | 0 | 50.0405 | 0.025 | 0.02497 | 0.55 | 249.7 | -0.4343 | 50.053 | 0.004508 | 0.13641 | 0.05439 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.015482 | 4.67 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 50.053 | 2.1994
2015-02-03 | 14:41 | 3505819312 | 0.012499 | 1.113 | 1001.44 | 17.1496 | 321.188 | 0 | 501.88 | 37.2202 | 1.113 | 0 | 25.0004 | 0.05 | 0.04997 | 0.57 | 499.7 | -0.4344 | 25.0129 | 0.012156 | 0.6952 | 0.13852 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.01 | 2.3591 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 25.0129 | 2.1996
2015-02-03 | 14:41 | 3505819312 | 0.012499 | 1.113 | 1001.44 | 17.1496 | 321.188 | 0 | 503.309 | 25.9224 | 1.113 | 0 | 25.0004 | 0.05 | 0.04997 | 0.55 | 499.7 | -0.7171 | 25.0129 | 0.010116 | 0.7491 | 0.14884 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.01 | 2.3591 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 25.0129 | 2.1996
2015-02-03 | 14:22 | 3505818131 | 0.024989 | 2.2386 | 1001.441 | 17.2146 | 321.135 | 0 | 999.727 | 52.9537 | 2.2394 | 0 | 24.983 | 0.1 | 0.09992 | 0.57 | 999.2 | -0.05271 | 25.008 | 0.007279 | 0.09396 | 0.009399 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 2.3568 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 25.008 | 2.1996
2015-02-03 | 14:22 | 3505818131 | 0.024989 | 2.2386 | 1001.441 | 17.2146 | 321.135 | 0 | 1001.853 | 16.0998 | 2.2394 | 0 | 24.983 | 0.1 | 0.09992 | 0.55 | 999.2 | -0.26481 | 25.008 | 0.004005 | 0.509 | 0.05081 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 2.3568 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 25.008 | 2.1996
2015-02-03 | 14:39 | 3505819146 | 0.025006 | 2.241 | 1001.44 | 17.131 | 351.069 | 0 | 500.635 | 12.3572 | 2.2417 | 0 | 50.0359 | 0.05 | 0.04995 | 0.57 | 499.5 | -0.22671 | 50.0609 | 0.007022 | 0.04539 | 0.009066 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.013609 | 4.6699 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 50.0609 | 2.1995
2015-02-03 | 14:39 | 3505819146 | 0.025006 | 2.241 | 1001.44 | 17.131 | 351.069 | 0 | 502.25 | 4.532 | 2.2417 | 0 | 50.0359 | 0.05 | 0.04995 | 0.55 | 499.5 | -0.5475 | 50.0609 | 0.004239 | 0.26495 | 0.05275 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.013609 | 4.6699 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 50.0609 | 2.1995
2015-02-03 | 14:19 | 3505817968 | 0.0375 | 3.374 | 1001.441 | 17.2422 | 334.248 | 0 | 997.703 | 39.9302 | 3.3744 | 0 | 37.4964 | 0.1 | 0.09991 | 0.57 | 999.1 | 0.14002 | 37.5339 | 0.006334 | 0.08159 | 0.008178 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 3.5518 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 37.5339 | 2.1993
2015-02-03 | 14:19 | 3505817968 | 0.0375 | 3.374 | 1001.441 | 17.2422 | 334.248 | 0 | 999.266 | 5.6637 | 3.3744 | 0 | 37.4964 | 0.1 | 0.09991 | 0.55 | 999.1 | -0.016612 | 37.5339 | 0.0023816 | 0.17875 | 0.017888 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 3.5518 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 37.5339 | 2.1993
2015-02-03 | 14:16 | 3505817805 | 0.05003 | 4.4728 | 1001.178 | 17.1851 | 350.585 | 0 | 999.06 | 41.3587 | 4.4728 | 0 | 50.0062 | 0.1 | 0.09994 | 0.57 | 999.4 | 0.03403 | 50.0562 | 0.006437 | 0.08304 | 0.008312 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 4.6654 | FF_Ch0_03-02-2015_4.xls | 6 | 1 | FF [C] | 50.0562 | 2.1996
2015-02-03 | 14:16 | 3505817805 | 0.05003 | 4.4728 | 1001.178 | 17.1851 | 350.585 | 0 | 1000.199 | 6.0849 | 4.4728 | 0 | 50.0062 | 0.1 | 0.09994 | 0.55 | 999.4 | -0.07988 | 50.0562 | 0.0024663 | 0.20376 | 0.020372 | 14:08 on 03-02-2015 | Trace_50_ FID_Calib_25.xls | 0.008 | 4.6654 | FF_Ch1_03-02-2015_4.xls | 6 | 1 | FF [C1] | 50.0562 | 2.1996

## Carrier1-50-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-02-02 | 9:00 | 3505712459 | 49.9853 | 4.6044 | 999.355 | 16.8758 | 353.138 | 296.76 | 246.58 | 49.8544 | 0.99738 |  | 62518.35 | 62764.93 | 49.8401 | 49.6225 | -0.028641 | -0.46514
2015-02-02 | 9:06 | 3505712798 | 40.0369 | 3.7225 | 999.355 | 16.9425 | 336.706 | 296.75 | 197.17 | 39.8659 | 0.99573 |  | 62782.3 | 62979.47 | 39.9091 | 40.1378 | 0.10834 | 0.68198
2015-02-02 | 9:11 | 3505713119 | 30.0245 | 2.7893 | 999.355 | 16.9949 | 324.562 | 296.75 | 148.24 | 29.9727 | 0.99827 |  | 62991.79 | 63140.03 | 29.9374 | 30.1013 | -0.11765 | 0.4289
2015-02-02 | 9:17 | 3505713440 | 20.0243 | 1.8388 | 999.355 | 17.0475 | 315.0693 | 296.76 | 98.87 | 19.9899 | 0.99828 |  | 63148.6 | 63247.47 | 19.974 | 19.8787 | -0.079407 | -0.55625
2015-02-02 | 9:22 | 3505713762 | 9.9911 | 0.8999 | 999.355 | 17.1244 | 307.6837 | 296.74 | 48.98 | 9.9036 | 0.99124 |  | 63252.32 | 63301.3 | 9.9446 | 9.7809 | 0.41343 | -1.2392
2015-02-02 | 9:28 | 3505714090 | 5.0047 | 0.4481 | 999.355 | 17.1989 | 303.747 | 296.73 | 24.43 | 4.9398 | 0.98704 |  | 63304.3 | 63328.73 | 4.9406 | 4.9218 | 0.015428 | -0.365
2015-02-02 | 9:33 | 3505714423 | 2.6657 | 0.24074 | 999.355 | 17.2839 | 300.7517 | 296.75 | 12.9 | 2.6083 | 0.97845 |  | 63330.5 | 63343.4 | 2.5888 | 2.6917 | -0.74717 | 3.1981
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | Low
[12 empty] | names | Air
[12 empty] | A | -0.19257 | 0.10253
[12 empty] | Bx | 11.6751 | 10.7549
[12 empty] | Cx^2 | -0.52409
[12 empty] | Dx^3 | 0.075673
[12 empty] | Calib Temp Deg C | 17.0668
[12 empty] | Calib Press Mbars | 999.355
[12 empty] | Date | 2015-02-02
Using domestic gas meter with a 1m long 22mm cooper pipe upstream to avoid jetting into meter

## Carrier1-50-CO2
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Calculated | Gas Factor | Calculated |  | voltage span fraction | Air:CO2 factor
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Carrier |  | Carrier |  | (%) | from cubic fit to Hi-Tec calcs |  | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | using gas meter |  | Flowrate |  |  |  |  | fit | fit |  | cubic | linear
Air
2015-02-02 | 9:00 | 3505712459 | 49.9853 | 4.6044 | 999.355 | 16.8758 | 353.138 | 49.8544 | 0.7429 | 37.037 | 0.74096 | 92.088 | 1.3461 |  | 37.0251 | 36.7349 |  | -0.032153 | -0.81582
2015-02-02 | 9:06 | 3505712798 | 40.0369 | 3.7225 | 999.355 | 16.9425 | 336.706 | 39.8659 | 0.73849 | 29.4407 | 0.73534 | 74.45 | 1.3541 |  | 29.4765 | 29.7032 |  | 0.12183 | 0.89164
2015-02-02 | 9:11 | 3505713119 | 30.0245 | 2.7893 | 999.355 | 16.9949 | 324.562 | 29.9727 | 0.73611 | 22.0632 | 0.73484 | 55.786 | 1.3585 |  | 22.0341 | 22.2624 |  | -0.13203 | 0.90287
2015-02-02 | 9:17 | 3505713440 | 20.0243 | 1.8388 | 999.355 | 17.0475 | 315.0693 | 19.9899 | 0.73554 | 14.7034 | 0.73428 | 36.776 | 1.3595 |  | 14.69 | 14.6837 |  | -0.091435 | -0.13394
2015-02-02 | 9:22 | 3505713762 | 9.9911 | 0.8999 | 999.355 | 17.1244 | 307.6837 | 9.9036 | 0.7363 | 7.292 | 0.72985 | 17.998 | 1.3581 |  | 7.3258 | 7.1976 |  | 0.46317 | -1.2957
2015-02-02 | 9:28 | 3505714090 | 5.0047 | 0.4481 | 999.355 | 17.1989 | 303.747 | 4.9398 | 0.73698 | 3.6406 | 0.72743 | 8.962 | 1.3569 |  | 3.6423 | 3.5952 |  | 0.04707 | -1.2464
2015-02-02 | 9:33 | 3505714423 | 2.6657 | 0.24074 | 999.355 | 17.2839 | 300.7517 | 2.6083 | 0.73734 | 1.9232 | 0.72145 | 4.8148 | 1.3562 |  | 1.9062 | 1.9418 |  | -0.88002 | 0.97054
[13 empty] | cubic | linear |  | Hi-Tec calculator output for F-201AC |  |  | cubic | linear
[16 empty] | x | y
[12 empty] | Hi-tec | Carrier 1 |  |  | 10 | 1.357 |  | 1.3554 | 1.3627 | A
[12 empty] | Calibration | Low |  |  | 25 | 1.359 |  | 0.00017346 | -0.00016792 | Bx
[12 empty] | names | CO2 |  |  | 50 | 1.359 |  | -7.9187e-07 |  | Cx^2
[12 empty] | A | -0.15177 | 0.022325 |  | 75 | 1.354 |  | -2.3823e-08 |  | Dx^3
[12 empty] | Bx | 8.651 | 7.9734 |  | 100 | 1.341
[12 empty] | Cx^2 | -0.44146
[12 empty] | Dx^3 | 0.068671
[12 empty] | Calib Temp Deg C | 17.0307
[12 empty] | Calib Press Mbars | 999.355
[12 empty] | Date | 2015-02-02

## Carrier1-5-Air Check Not used
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-02-02 | 14:49 | 3505733395 | 4.9726 | 4.572 | 998.461 | 16.8258 | 306.8154 |  |  | 4.965 | 0.99847 | bubble flow 0.02 - 6l/min |  |  | 4.9645 | 4.944 | -0.0094662 | -0.42325
2015-02-02 | 14:52 | 3505733571 | 4.0019 | 3.7218 | 998.461 | 16.5674 | 305.4964 |  |  | 3.986 | 0.99603 | bubble flow 0.02 - 6l/min |  |  | 3.9991 | 4.0272 | 0.32769 | 1.0327
2015-02-02 | 16:42 | 3505740128 | 2.9995 | 2.783 | 998.263 | 17.373 | 304.8731 |  |  | 3.026 | 1.0088 | bubble flow 0.02 - 6l/min |  |  | 2.9885 | 3.0148 | -1.2405 | -0.37008
2015-02-02 | 16:44 | 3505740273 | 2.0006 | 1.8375 | 998.263 | 17.2935 | 302.8782 |  |  | 2.006 | 1.0027 | bubble flow 0.02 - 6l/min |  |  | 1.9938 | 1.9952 | -0.60826 | -0.53774
2015-02-02 | 16:47 | 3505740436 | 0.9988 | 0.9013 | 998.263 | 17.2943 | 300.2963 |  |  | 0.9958 | 0.997 | bubble flow 0.02 - 6l/min |  |  | 0.99538 | 0.98565 | -0.04205 | -1.019
2015-02-02 | 16:49 | 3505740596 | 0.5 | 0.4497 | 998.263 | 17.3925 | 298.6494 |  |  | 0.4974 | 0.9948 | bubble flow 0.02 - 6l/min |  |  | 0.49824 | 0.49867 | 0.16788 | 0.25457
2015-02-02 | 16:52 | 3505740779 | 0.24852 | 0.227 | 998.263 | 17.4483 | 297.3428 |  |  | 0.2478 | 0.9971 | bubble flow 0.02 - 6l/min |  |  | 0.24758 | 0.25852 | -0.088883 | 4.3243
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | Very Low
[12 empty] | names | Air
[12 empty] | A | -0.012325 | 0.013728
[12 empty] | Bx | 1.1555 | 1.0784
[12 empty] | Cx^2 | -0.048158
[12 empty] | Dx^3 | 0.00733
[12 empty] | Calib Temp Deg C | 17.1707
[12 empty] | Calib Press Mbars | 998.3196
[12 empty] | Date | 2015-02-02

## Carrier1-5-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-01-20 | 11:20 | 3504597623 | 5.0015 | 4.5744 | 1004.32 | 17.7411 | 1.7336 | 296.73 | 24.55 | 4.9641 | 0.99252 |  | 57072.8 | 57097.35 | 4.9673 | 4.9404 | 0.06529 | -0.47672
2015-01-20 | 11:26 | 3504597968 | 4.0053 | 3.7168 | 1003.923 | 17.7591 | 1.0803 | 296.72 | 19.8 | 4.0038 | 0.99962 |  | 57100.6 | 57120.4 | 3.9936 | 4.0161 | -0.25527 | 0.30672
2015-01-20 | 11:32 | 3504598320 | 3.0021 | 2.8001 | 1003.923 | 17.7787 | 0.485 | 296.73 | 14.82 | 2.9967 | 0.99819 |  | 57123.23 | 57138.05 | 3.0066 | 3.028 | 0.33009 | 1.0445
2015-01-20 | 11:38 | 3504598690 | 2.0015 | 1.8549 | 1003.923 | 17.7978 | 0.1004 | 296.74 | 9.95 | 2.0119 | 1.0052 |  | 57140.7 | 57150.65 | 2.0121 | 2.0092 | 0.012524 | -0.13467
2015-01-20 | 13:31 | 3504605496 | 0.999 | 0.9112 | 1002.831 | 18.0205 | -0.5529 | 116.7 |  | 1.013 | 1.014 | bubble flow 0.02 - 6l/min | 57152.8 | 57157.6 | 1.0061 | 0.99196 | -0.67764 | -2.0771
2015-01-20 | 13:34 | 3504605656 | 0.5988 | 0.535 | 1002.831 | 17.9965 | -0.7944 | 116.7 |  | 0.592 | 0.98864 | bubble flow 0.02 - 6l/min | 57159.55 | 57163.35 | 0.59321 | 0.58646 | 0.20452 | -0.93566
2015-01-20 | 13:04 | 3504603848 | 0.26968 | 0.227 | 1003.327 | 17.9725 | -0.8748 | 116.78 |  | 0.2451 | 0.90885 | bubble flow 0.02 - 6l/min |  |  | 0.24758 | 0.25447 | 1.0117 | 3.8247
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | Very Low
[12 empty] | names | Air
[12 empty] | A | -0.012325 | 0.0097958
[12 empty] | Bx | 1.1555 | 1.0779
[12 empty] | Cx^2 | -0.048158
[12 empty] | Dx^3 | 0.00733
[12 empty] | Calib Temp Deg C | 17.8666
[12 empty] | Calib Press Mbars | 1003.5826
[12 empty] | Date | 2015-01-20

## Carrier2-50-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 2 | Carrier 2 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-02-02 | 11:12 | 3505720324 | 49.9743 | 4.6457 | 999.455 | 17.3241 | 355.449 | 296.72 | 246.84 | 49.9137 | 0.99879 |  | 64196.2 | 64443.04 | 49.904 | 49.578 | -0.019495 | -0.67256
2015-02-02 | 11:17 | 3505720645 | 40.027 | 3.7712 | 999.454 | 17.3539 | 338.117 | 296.7 | 197.55 | 39.9494 | 0.99806 |  | 64459.3 | 64656.85 | 39.979 | 40.2523 | 0.074063 | 0.7581
2015-02-02 | 11:22 | 3505720964 | 30.0141 | 2.8331 | 999.455 | 17.3655 | 325.752 | 296.7 | 148.43 | 30.0162 | 1.0001 |  | 64668.62 | 64817.05 | 29.9916 | 30.2483 | -0.081985 | 0.77347
2015-02-02 | 11:28 | 3505721285 | 19.9993 | 1.8681 | 999.454 | 17.4009 | 316.026 | 296.71 | 98.95 | 20.0094 | 1.0005 |  | 64825.55 | 64924.5 | 19.9984 | 19.9575 | -0.055218 | -0.25945
2015-02-02 | 11:33 | 3505721607 | 10.0134 | 0.9206 | 999.454 | 17.4646 | 308.4682 | 296.7 | 49.4 | 9.9899 | 0.99765 |  | 64929.3 | 64978.7 | 10.0265 | 9.8533 | 0.36655 | -1.367
2015-02-02 | 11:38 | 3505721938 | 5.018 | 0.464 | 999.455 | 17.5331 | 304.5602 | 296.7 | 25 | 5.0556 | 1.0075 |  | 64981.9 | 65006.9 | 5.036 | 4.9841 | -0.38743 | -1.4143
2015-02-02 | 11:44 | 3505722270 | 2.6883 | 0.25642 | 999.455 | 17.6009 | 301.6031 | 296.7 | 13.4 | 2.7098 | 1.008 |  | 65008.62 | 65022.02 | 2.7087 | 2.7705 | -0.042266 | 2.2384
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 2
[12 empty] | Calibration | Low
[12 empty] | names | Air
[12 empty] | A | -0.22611 | 0.035986
[12 empty] | Bx | 11.5842 | 10.6641
[12 empty] | Cx^2 | -0.56371
[12 empty] | Dx^3 | 0.084573
[12 empty] | Calib Temp Deg C | 17.4347
[12 empty] | Calib Press Mbars | 999.4546
[12 empty] | Date | 2015-02-02
Using domestic gas meter with a 1m long 22mm cooper pipe upstream to avoid jetting into meter