## Carrier1-500-CO2
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Calculated | Gas Factor | Calculated |  | voltage span fraction | Air:CO2 factor
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Carrier |  | Carrier |  | (%) | from cubic fit to Hi-Tec calcs |  | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Flowrate Using FFID |  | Flowrate Using FFID |  |  |  |  | fit | fit |  | cubic | linear
Air
2015-03-16 | 14:26 | 3509360797 | 49.6541 | 0.27842 | 1014.945 | 18.9162 | 470.81 | 50.619 | 0.73564 | 37.2376 | 0.74994 | 5.5684 | 1.3594 |  | 37.4113 | 39.6736 |  | 0.46639 | 6.5419
2015-03-16 | 14:28 | 3509360893 | 99.4031 | 0.5497 | 1014.945 | 18.9187 | 1451.547 | 101.1001 | 0.73519 | 74.3277 | 0.74774 | 10.994 | 1.3602 |  | 74.0241 | 74.057 |  | -0.40837 | -0.36413
2015-03-16 | 14:29 | 3509360984 | 199.6915 | 1.1117 | 1014.945 | 18.9212 | 653.6 | 200.5779 | 0.73442 | 147.3085 | 0.73768 | 22.234 | 1.3616 |  | 147.4073 | 145.2877 |  | 0.067082 | -1.3718
2015-03-16 | 14:31 | 3509361070 | 299.7041 | 1.6887 | 1014.945 | 18.9515 | 1199.538 | 299.8417 | 0.73397 | 220.075 | 0.73431 | 33.774 | 1.3625 |  | 220.24 | 218.4196 |  | 0.074982 | -0.7522
2015-03-16 | 14:33 | 3509361193 | 399.817 | 2.2766 | 1014.945 | 18.9771 | 1272.159 | 399.52 | 0.73399 | 293.2424 | 0.73344 | 45.532 | 1.3624 |  | 293.049 | 292.933 |  | -0.065926 | -0.1055
2015-03-16 | 14:34 | 3509361261 | 499.696 | 2.8662 | 1014.945 | 18.9586 | 1748.383 | 498.009 | 0.73461 | 365.8418 | 0.73213 | 57.324 | 1.3613 |  | 365.9011 | 367.6619 |  | 0.016219 | 0.49751
[13 empty] | cubic | linear |  | Hi-Tec calculator output for F-201AC |  |  | cubic | linear
[16 empty] | x | y
[12 empty] | Hi-tec | Carrier 1 |  |  | 10 | 1.36 |  | 1.3584 | 1.3657 | A
[12 empty] | Calibration | High |  |  | 25 | 1.362 |  | 0.00017346 | -0.00016792 | Bx
[12 empty] | names | CO2 |  |  | 50 | 1.362 |  | -7.9187e-07 |  | Cx^2
[12 empty] | A | -1.1461 | 4.3853 |  | 75 | 1.357 |  | -2.3823e-08 |  | Dx^3
[12 empty] | Bx | 140.4277 | 126.745 |  | 100 | 1.344
[12 empty] | Cx^2 | -7.2593
[12 empty] | Dx^3 | 1.0273
[12 empty] | Calib Temp Deg C | 18.9405
[12 empty] | Calib Press Mbars | 1014.945
[12 empty] | Date | 2015-03-16

## Carrier1-500-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU |  |  | Calculated
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source |  |  | Carrier | [5 empty] | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure |  |  | Flowrate Using FFID | [5 empty] | fit | fit |  | cubic | linear
2015-03-16 | 14:26 | 3509360797 | 49.6541 | 0.27842 | 1014.945 | 18.9162 | 470.81 |  |  | 50.619 | 1.0194 |  |  |  |  | 50.8598 | 53.9784 |  | 0.47563 | 6.6367
2015-03-16 | 14:28 | 3509360893 | 99.4031 | 0.5497 | 1014.945 | 18.9187 | 1451.547 |  |  | 101.1001 | 1.0171 |  |  |  |  | 100.6798 | 100.809 |  | -0.41572 | -0.28797
2015-03-16 | 14:29 | 3509360984 | 199.6915 | 1.1117 | 1014.945 | 18.9212 | 653.6 |  |  | 200.5779 | 1.0044 |  |  |  |  | 200.713 | 197.8259 |  | 0.067365 | -1.372
2015-03-16 | 14:31 | 3509361070 | 299.7041 | 1.6887 | 1014.945 | 18.9515 | 1199.538 |  |  | 299.8417 | 1.0005 |  |  |  |  | 300.0732 | 297.4323 |  | 0.077203 | -0.80356
2015-03-16 | 14:33 | 3509361193 | 399.817 | 2.2766 | 1014.945 | 18.9771 | 1272.159 |  |  | 399.52 | 0.99926 |  |  |  |  | 399.2502 | 398.9203 |  | -0.067523 | -0.1501
2015-03-16 | 14:34 | 3509361261 | 499.696 | 2.8662 | 1014.945 | 18.9586 | 1748.383 |  |  | 498.009 | 0.99662 |  |  |  |  | 498.0917 | 500.7018 |  | 0.016609 | 0.54071
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | High
[12 empty] | names | Air
[12 empty] | A | -1.5146 | 5.9153
[12 empty] | Bx | 190.5584 | 172.628
[12 empty] | Cx^2 | -9.1189
[12 empty] | Dx^3 | 1.2036
[12 empty] | Calib Temp Deg C | 18.9405
[12 empty] | Calib Press Mbars | 1014.945
[12 empty] | Date | 2015-03-16
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Qtrace | FF (C) | FF (C) | GMU | GMU | Calculated | GMU | GMU | FF (C) | Cppm | Hi-Tec/FID | Q | c'/C | FF (C) | C | Multi Run | File | GMU | GMU | FF (C) | No Of | Flow | Inst | GMU | Clean_Air | Calculated
of | Measurement | since | Trace | Trace | NPL | Therm | Source | From | Mean | Variance | Trace | Carrier 1 | Carrier | Source | Source | Calibration | Hi-Tec | % | Source |  | Standard Error | % | Start Time |  | Carrier 1 | Carrier 2 | Calibration | Measurement | Measurement |  | Source | Hi-Tec | Carrier
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | FID | C | c^2 | Dem Volt | Meas Q | Flowrate | Dem Strength | Meas Strength | Drift |  |  | l/min |  | in Mean | Error |  |  | Meas Volt | Meas Volt | Filename | Blocks | Time |  | Meas Q | Meas Volt | Flowrate
2015-03-16 | 14:26 | 3509360797 | 0.024992 | 2.2565 | 1014.945 | 18.9162 | 470.81 | 0.024516 | 493.728 | 36.9886 | 2.2573 | 49.6541 | 50.619 | 0.05 | 0.05031 | 0.12 | 503.1 | 1.8982 | 49.6791 | 0.012318 | 1.5112 | 0.30608 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 0.27842 | 0.011009 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 49.6791 | 1.9966 | 50.619
2015-03-16 | 14:28 | 3509360893 | 0.05003 | 4.506 | 1014.945 | 18.9187 | 1451.547 | 0.04919 | 494.856 | 33.8346 | 4.5034 | 99.4031 | 101.1001 | 0.05 | 0.05031 | 0.12 | 503.1 | 1.6659 | 99.4531 | 0.011754 | 1.6775 | 0.339 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 0.5497 | 0.007 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 99.4531 | 1.9978 | 101.1001
2015-03-16 | 14:29 | 3509360984 | 0.05003 | 4.5061 | 1014.945 | 18.9212 | 653.6 | 0.04981 | 249.4293 | 16.3581 | 4.5034 | 199.6915 | 200.5779 | 0.025 | 0.025048 | 0.12 | 250.48 | 0.4212 | 199.7415 | 0.016215 | 0.837 | 0.3356 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 1.1117 | 0.005474 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 199.7415 | 1.9982 | 200.5779
2015-03-16 | 14:31 | 3509361070 | 0.04502 | 4.071 | 1014.945 | 18.9515 | 1199.538 | 0.045 | 150.1459 | 2.5672 | 4.0697 | 299.7041 | 299.8417 | 0.015 | 0.015018 | 0.12 | 150.18 | 0.022711 | 299.7491 | 0.010671 | 0.3389 | 0.22571 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 1.6887 | 0.004721 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 299.7491 | 1.998 | 299.8417
2015-03-16 | 14:33 | 3509361193 | 0.05202 | 4.675 | 1014.945 | 18.9771 | 1272.159 | 0.05206 | 130.2062 | 3.7283 | 4.6736 | 399.817 | 399.52 | 0.013 | 0.013009 | 0.12 | 130.09 | -0.08924 | 399.869 | 0.014829 | 0.3714 | 0.28524 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 2.2766 | 0.004 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 399.869 | 1.9978 | 399.52
2015-03-16 | 14:34 | 3509361261 | 0.05004 | 4.5061 | 1014.945 | 18.9586 | 1748.383 | 0.05021 | 100.4801 | 2.3572 | 4.5032 | 499.696 | 498.009 | 0.01 | 0.010012 | 0.12 | 100.12 | -0.3584 | 499.746 | 0.01528 | 0.3256 | 0.324 | 14:18 on 16-03-2015 | Carrier_500_FID_Calib_1.xls | 2.8662 | 0.00384 | FF_Ch0_16-03-2015_4.xls | 1 | 0.5 |  | 499.746 | 1.9976 | 498.009

## Trace 300
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Trace | Trace | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-01-26 | 12:43 | 3505120982 | 0.25003 | 4.1576 | 1015.839 | 17.835 | 344.341 |  |  | 0.2516 | 1.0063 | bubble flow 1-250cc |  |  | 0.25155 | 0.25125 | -0.019723 | -0.14029
2015-01-26 | 12:45 | 3505121147 | 0.19988 | 3.3132 | 1015.839 | 17.7434 | 344.243 |  |  | 0.1996 | 0.9986 | bubble flow 1-250cc |  |  | 0.19976 | 0.2006 | 0.078841 | 0.50188
2015-01-26 | 12:48 | 3505121309 | 0.1501 | 2.4515 | 1015.839 | 17.8484 | 344.175 |  |  | 0.149 | 0.99267 | bubble flow 1-250cc |  |  | 0.14882 | 0.14892 | -0.11813 | -0.054468
2015-01-26 | 12:50 | 3505121430 | 0.09995 | 1.5898 | 1015.996 | 17.732 | 344.139 |  |  | 0.09817 | 0.98219 | bubble flow 1-250cc |  |  | 0.098187 | 0.097236 | 0.017631 | -0.95147
2015-01-26 | 12:52 | 3505121545 | 0.05008 | 0.7591 | 1016.137 | 17.8065 | 344.218 |  |  | 0.048 | 0.95847 | bubble flow 1-250cc |  |  | 0.048106 | 0.047412 | 0.21994 | -1.2243
2015-01-26 | 12:54 | 3505121690 | 0.024987 | 0.3566 | 1016.137 | 17.8 | 344.238 |  |  | 0.02295 | 0.91848 | bubble flow 1-250cc |  |  | 0.022944 | 0.023271 | -0.026123 | 1.3999
2015-01-26 | 12:57 | 3505121867 | 0.018716 | 0.25909 | 1016.137 | 17.8756 | 344.186 |  |  | 0.01679 | 0.89709 | bubble flow 1-250cc |  |  | 0.016732 | 0.017423 | -0.3429 | 3.7691
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Trace
[12 empty] | Calibration | 300 ml/min
[12 empty] | names | Propane
[12 empty] | A | -2.2e-05 | 0.0018832
[12 empty] | Bx | 0.065407 | 0.059978
[12 empty] | Cx^2 | -0.00297
[12 empty] | Dx^3 | 0.000431
[12 empty] | Calib Temp Deg C | 17.8058
[12 empty] | Calib Press Mbars | 1015.9891
[12 empty] | Date | 2015-01-26

## Trace 50
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Bubble |  |  | Start | End
of | Measurement | since | Trace | Trace | NPL | Therm | Source | Flowrate |  |  | Gas | Gas |  | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure |  |  |  | Reading | Reading |  | fit | fit | cubic | linear
2016-06-15 |  |  | 0.045 | 3.9312 | 990 | 21.6 | 1 | 0.04351 | 0.96689 | bubble flow 1-250cc |  |  |  | 0.043883 | 0.044048 | 0.85616 | 1.2357
2016-06-15 |  |  | 0.0525 | 4.5645 | 990 | 21.6 | 1 | 0.05138 | 0.97867 | bubble flow 1-250cc |  |  |  | 0.05124 | 0.051125 | -0.27283 | -0.49716
2016-06-15 |  |  | 0.0375 | 3.2858 | 990 | 21.6 | 1 | 0.03691 | 0.98427 | bubble flow 1-250cc |  |  |  | 0.036676 | 0.036836 | -0.63467 | -0.2017
2016-06-15 |  |  | 0.03 | 2.6238 | 990 | 21.6 | 1 | 0.02954 | 0.98467 | bubble flow 1-250cc |  |  |  | 0.029445 | 0.029438 | -0.32141 | -0.34548
2016-06-15 |  |  | 0.0225 | 1.9575 | 990 | 21.6 | 1 | 0.02215 | 0.98444 | bubble flow 1-250cc |  |  |  | 0.022183 | 0.021992 | 0.14853 | -0.71203
2016-06-15 |  |  | 0.015 | 1.2953 | 990 | 21.6 | 1 | 0.01475 | 0.98333 | bubble flow 1-250cc |  |  |  | 0.014831 | 0.014592 | 0.54829 | -1.0682
2016-06-15 |  |  | 0.0075 | 0.6453 | 990 | 21.6 | 1 | 0.007301 | 0.97347 | bubble flow 1-250cc |  |  |  | 0.0073393 | 0.0073289 | 0.52484 | 0.38251
2016-06-15 |  |  | 0.0037 | 0.3275 | 990 | 21.6 | 1 | 0.003596 | 0.97189 | bubble flow 1-250cc |  |  |  | 0.0035344 | 0.0037776 | -1.7121 | 5.0508
[11 empty] | cubic | linear
[10 empty] | Hi-tec | Trace
[10 empty] | Calibration | 50 ml/min
[10 empty] | names | Propane
[10 empty] | A | -0.00050766 | 0.00011793
[10 empty] | Bx | 0.012548 | 0.011175
[10 empty] | Cx^2 | -0.00065622
[10 empty] | Dx^3 | 8.5639e-05
[10 empty] | Calib Temp Deg C | 21.6
[10 empty] | Calib Press Mbars | 990
[10 empty] | Date | 2016-06-15 |  | See correction below
0.001

## Trace 25
Date | Output | GMU | Customer flow | Customer flow
of | signal | Trace | C3H8 | C3H8
Measurement | % | Meas Volt | ml/min | l/min
2017-01-30 | 100.21 | 5.0105 | 25.05 | 0.02505
2017-01-30 | 49.73 | 2.4865 | 12.43 | 0.01243
2017-01-30 | 0 | 0 | 0 | 0
[5 empty] | Hi-tec | Trace
[5 empty] | Calibration | 25 ml/min
[5 empty] | names | Propane
[5 empty] | A | -4e-07
[5 empty] | Bx | 0.0049995
[5 empty] | Cx^2 | 0
[5 empty] | Dx^3 | 0
[5 empty] | Calib Temp Deg C | 0
[5 empty] | Calib Press Mbars | 1013.25
[5 empty] | Date | 2017-01-30

## Carrier1-50-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-02-02 | 9:00 | 3505712459 | 49.9853 | 4.6044 | 999.355 | 16.8758 | 353.138 | 296.76 | 246.58 | 49.8544 | 0.99738 |  | 62518.35 | 62764.93 | 49.8401 | 49.6225 | -0.028641 | -0.46514
2015-02-02 | 9:06 | 3505712798 | 40.0369 | 3.7225 | 999.355 | 16.9425 | 336.706 | 296.75 | 197.17 | 39.8659 | 0.99573 |  | 62782.3 | 62979.47 | 39.9091 | 40.1378 | 0.10834 | 0.68198
2015-02-02 | 9:11 | 3505713119 | 30.0245 | 2.7893 | 999.355 | 16.9949 | 324.562 | 296.75 | 148.24 | 29.9727 | 0.99827 |  | 62991.79 | 63140.03 | 29.9374 | 30.1013 | -0.11765 | 0.4289
2015-02-02 | 9:17 | 3505713440 | 20.0243 | 1.8388 | 999.355 | 17.0475 | 315.0693 | 296.76 | 98.87 | 19.9899 | 0.99828 |  | 63148.6 | 63247.47 | 19.974 | 19.8787 | -0.079407 | -0.55625
2015-02-02 | 9:22 | 3505713762 | 9.9911 | 0.8999 | 999.355 | 17.1244 | 307.6837 | 296.74 | 48.98 | 9.9036 | 0.99124 |  | 63252.32 | 63301.3 | 9.9446 | 9.7809 | 0.41343 | -1.2392
2015-02-02 | 9:28 | 3505714090 | 5.0047 | 0.4481 | 999.355 | 17.1989 | 303.747 | 296.73 | 24.43 | 4.9398 | 0.98704 |  | 63304.3 | 63328.73 | 4.9406 | 4.9218 | 0.015428 | -0.365
2015-02-02 | 9:33 | 3505714423 | 2.6657 | 0.24074 | 999.355 | 17.2839 | 300.7517 | 296.75 | 12.9 | 2.6083 | 0.97845 |  | 63330.5 | 63343.4 | 2.5888 | 2.6917 | -0.74717 | 3.1981
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | Low
[12 empty] | names | Air
[12 empty] | A | -0.19257 | 0.10253
[12 empty] | Bx | 11.6751 | 10.7549
[12 empty] | Cx^2 | -0.52409
[12 empty] | Dx^3 | 0.075673
[12 empty] | Calib Temp Deg C | 17.0668
[12 empty] | Calib Press Mbars | 999.355
[12 empty] | Date | 2015-02-02
Using domestic gas meter with a 1m long 22mm cooper pipe upstream to avoid jetting into meter

## Carrier1-50-CO2
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Calculated | Gas Factor | Calculated |  | voltage span fraction | Air:CO2 factor
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Carrier |  | Carrier |  | (%) | from cubic fit to Hi-Tec calcs |  | cubit | linear |  | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | using gas meter |  | Flowrate |  |  |  |  | fit | fit |  | cubic | linear
Air
2015-02-02 | 9:00 | 3505712459 | 49.9853 | 4.6044 | 999.355 | 16.8758 | 353.138 | 49.8544 | 0.7429 | 37.037 | 0.74096 | 92.088 | 1.3461 |  | 37.0251 | 36.7349 |  | -0.032153 | -0.81582
2015-02-02 | 9:06 | 3505712798 | 40.0369 | 3.7225 | 999.355 | 16.9425 | 336.706 | 39.8659 | 0.73849 | 29.4407 | 0.73534 | 74.45 | 1.3541 |  | 29.4765 | 29.7032 |  | 0.12183 | 0.89164
2015-02-02 | 9:11 | 3505713119 | 30.0245 | 2.7893 | 999.355 | 16.9949 | 324.562 | 29.9727 | 0.73611 | 22.0632 | 0.73484 | 55.786 | 1.3585 |  | 22.0341 | 22.2624 |  | -0.13203 | 0.90287
2015-02-02 | 9:17 | 3505713440 | 20.0243 | 1.8388 | 999.355 | 17.0475 | 315.0693 | 19.9899 | 0.73554 | 14.7034 | 0.73428 | 36.776 | 1.3595 |  | 14.69 | 14.6837 |  | -0.091435 | -0.13394
2015-02-02 | 9:22 | 3505713762 | 9.9911 | 0.8999 | 999.355 | 17.1244 | 307.6837 | 9.9036 | 0.7363 | 7.292 | 0.72985 | 17.998 | 1.3581 |  | 7.3258 | 7.1976 |  | 0.46317 | -1.2957
2015-02-02 | 9:28 | 3505714090 | 5.0047 | 0.4481 | 999.355 | 17.1989 | 303.747 | 4.9398 | 0.73698 | 3.6406 | 0.72743 | 8.962 | 1.3569 |  | 3.6423 | 3.5952 |  | 0.04707 | -1.2464
2015-02-02 | 9:33 | 3505714423 | 2.6657 | 0.24074 | 999.355 | 17.2839 | 300.7517 | 2.6083 | 0.73734 | 1.9232 | 0.72145 | 4.8148 | 1.3562 |  | 1.9062 | 1.9418 |  | -0.88002 | 0.97054
[13 empty] | cubic | linear |  | Hi-Tec calculator output for F-201AC |  |  | cubic | linear
[16 empty] | x | y
[12 empty] | Hi-tec | Carrier 1 |  |  | 10 | 1.357 |  | 1.3554 | 1.3627 | A
[12 empty] | Calibration | Low |  |  | 25 | 1.359 |  | 0.00017346 | -0.00016792 | Bx
[12 empty] | names | CO2 |  |  | 50 | 1.359 |  | -7.9187e-07 |  | Cx^2
[12 empty] | A | -0.15177 | 0.022325 |  | 75 | 1.354 |  | -2.3823e-08 |  | Dx^3
[12 empty] | Bx | 8.651 | 7.9734 |  | 100 | 1.341
[12 empty] | Cx^2 | -0.44146
[12 empty] | Dx^3 | 0.068671
[12 empty] | Calib Temp Deg C | 17.0307
[12 empty] | Calib Press Mbars | 999.355
[12 empty] | Date | 2015-02-02

## Carrier1-5-Air Check Not used
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-02-02 | 14:49 | 3505733395 | 4.9726 | 4.572 | 998.461 | 16.8258 | 306.8154 |  |  | 4.965 | 0.99847 | bubble flow 0.02 - 6l/min |  |  | 4.9645 | 4.944 | -0.0094662 | -0.42325
2015-02-02 | 14:52 | 3505733571 | 4.0019 | 3.7218 | 998.461 | 16.5674 | 305.4964 |  |  | 3.986 | 0.99603 | bubble flow 0.02 - 6l/min |  |  | 3.9991 | 4.0272 | 0.32769 | 1.0327
2015-02-02 | 16:42 | 3505740128 | 2.9995 | 2.783 | 998.263 | 17.373 | 304.8731 |  |  | 3.026 | 1.0088 | bubble flow 0.02 - 6l/min |  |  | 2.9885 | 3.0148 | -1.2405 | -0.37008
2015-02-02 | 16:44 | 3505740273 | 2.0006 | 1.8375 | 998.263 | 17.2935 | 302.8782 |  |  | 2.006 | 1.0027 | bubble flow 0.02 - 6l/min |  |  | 1.9938 | 1.9952 | -0.60826 | -0.53774
2015-02-02 | 16:47 | 3505740436 | 0.9988 | 0.9013 | 998.263 | 17.2943 | 300.2963 |  |  | 0.9958 | 0.997 | bubble flow 0.02 - 6l/min |  |  | 0.99538 | 0.98565 | -0.04205 | -1.019
2015-02-02 | 16:49 | 3505740596 | 0.5 | 0.4497 | 998.263 | 17.3925 | 298.6494 |  |  | 0.4974 | 0.9948 | bubble flow 0.02 - 6l/min |  |  | 0.49824 | 0.49867 | 0.16788 | 0.25457
2015-02-02 | 16:52 | 3505740779 | 0.24852 | 0.227 | 998.263 | 17.4483 | 297.3428 |  |  | 0.2478 | 0.9971 | bubble flow 0.02 - 6l/min |  |  | 0.24758 | 0.25852 | -0.088883 | 4.3243
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | Very Low
[12 empty] | names | Air
[12 empty] | A | -0.012325 | 0.013728
[12 empty] | Bx | 1.1555 | 1.0784
[12 empty] | Cx^2 | -0.048158
[12 empty] | Dx^3 | 0.00733
[12 empty] | Calib Temp Deg C | 17.1707
[12 empty] | Calib Press Mbars | 998.3196
[12 empty] | Date | 2015-02-02

## Carrier1-5-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  | Bubble
of | Measurement | since | Carrier 1 | Carrier 1 | NPL | Therm | Source | Interval | Used | Flowrate |  |  |  |  | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds | [6 empty] | fit | fit | cubic | linear
2016-06-15 |  |  | 1.0002 | 0.8827 | 990.19 | 21.3 | 1 |  | 0 | 1.001 | 1.0008 | bubble flow 0.02 - 6l/min |  |  | 1.0006 | 0.98916 | -0.042759 | -1.1826
2016-06-15 |  |  | 1.4002 | 1.2449 | 990.19 | 21.3 | 1 |  | 0 | 1.403 | 1.002 | bubble flow 0.02 - 6l/min |  |  | 1.4023 | 1.3847 | -0.046368 | -1.3046
2016-06-15 |  |  | 2.0003 | 1.7965 | 990.19 | 21.3 | 1 |  | 0 | 2.002 | 1.0008 | bubble flow 0.02 - 6l/min |  |  | 2.0018 | 1.9871 | -0.0086476 | -0.74618
2016-06-15 |  |  | 3.0001 | 2.724 | 990.19 | 21.3 | 1 |  | 0 | 2.994 | 0.99797 | bubble flow 0.02 - 6l/min |  |  | 2.9937 | 2.9999 | -0.010929 | 0.19778
2016-06-15 |  |  | 4.0006 | 3.6337 | 990.19 | 21.3 | 1 |  |  | 3.974 | 0.99335 | bubble flow 0.02 - 6l/min |  |  | 3.9746 | 3.9933 | 0.01463 | 0.48675
2016-06-15 |  |  | 5.001 | 4.498 | 990.19 | 21.3 | 1 |  |  | 4.943 | 0.9884 | bubble flow 0.02 - 6l/min |  |  | 4.9428 | 4.9372 | -0.0042474 | -0.1176
2016-06-15 |  |  | 1.0003 | 0.8831 | 990.19 | 21.3 | 1 |  | 0 | 1 | 0.9997 | bubble flow 0.02 - 6l/min |  |  | 1.001 | 0.9896 | 0.10202 | -1.0401
2016-06-15 |  |  | 0.4999 | 0.4394 | 990.19 | 21.3 | 1 |  | 0 | 0.4955 | 0.9912 | bubble flow 0.02 - 6l/min |  |  | 0.49668 | 0.50506 | 0.23797 | 1.9302
2016-06-15 |  |  | 0.2001 | 0.1808 | 990.19 | 21.3 | 1 |  |  | 0.1962 | 0.98051 | bubble flow 0.02 - 6l/min |  |  | 0.19521 | 0.22266 | -0.50454 | 13.4887
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 1
[12 empty] | Calibration | Very Low
[12 empty] | names | Air
[12 empty] | A | -0.019342 | 0.025225
[12 empty] | Bx | 1.1958 | 1.092
[12 empty] | Cx^2 | -0.0519
[12 empty] | Dx^3 | 0.006959
[12 empty] | Calib Temp Deg C | 21.3
[12 empty] | Calib Press Mbars | 990.19
[12 empty] | Date | 2016-06-15

## Carrier2-50-Air
Date | Time | Seconds | GMU | GMU | Froude | EnFlo_Tunnel | GMU | Image |  |  |  |  | Start | End
of | Measurement | since | Carrier 2 | Carrier 2 | NPL | Therm | Source | Interval | Used | Flowrate |  |  | Gas | Gas | cubit | linear | % err. | % err.
Measurement | Commenced | 1st Jan | Meas Q | Meas Volt | Atmos_Press | Master_Reference | Pressure | Seconds |  |  |  |  | Reading | Reading | fit | fit | cubic | linear
2015-02-02 | 11:12 | 3505720324 | 49.9743 | 4.6457 | 999.455 | 17.3241 | 355.449 | 296.72 | 246.84 | 49.9137 | 0.99879 |  | 64196.2 | 64443.04 | 49.904 | 49.578 | -0.019495 | -0.67256
2015-02-02 | 11:17 | 3505720645 | 40.027 | 3.7712 | 999.454 | 17.3539 | 338.117 | 296.7 | 197.55 | 39.9494 | 0.99806 |  | 64459.3 | 64656.85 | 39.979 | 40.2523 | 0.074063 | 0.7581
2015-02-02 | 11:22 | 3505720964 | 30.0141 | 2.8331 | 999.455 | 17.3655 | 325.752 | 296.7 | 148.43 | 30.0162 | 1.0001 |  | 64668.62 | 64817.05 | 29.9916 | 30.2483 | -0.081985 | 0.77347
2015-02-02 | 11:28 | 3505721285 | 19.9993 | 1.8681 | 999.454 | 17.4009 | 316.026 | 296.71 | 98.95 | 20.0094 | 1.0005 |  | 64825.55 | 64924.5 | 19.9984 | 19.9575 | -0.055218 | -0.25945
2015-02-02 | 11:33 | 3505721607 | 10.0134 | 0.9206 | 999.454 | 17.4646 | 308.4682 | 296.7 | 49.4 | 9.9899 | 0.99765 |  | 64929.3 | 64978.7 | 10.0265 | 9.8533 | 0.36655 | -1.367
2015-02-02 | 11:38 | 3505721938 | 5.018 | 0.464 | 999.455 | 17.5331 | 304.5602 | 296.7 | 25 | 5.0556 | 1.0075 |  | 64981.9 | 65006.9 | 5.036 | 4.9841 | -0.38743 | -1.4143
2015-02-02 | 11:44 | 3505722270 | 2.6883 | 0.25642 | 999.455 | 17.6009 | 301.6031 | 296.7 | 13.4 | 2.7098 | 1.008 |  | 65008.62 | 65022.02 | 2.7087 | 2.7705 | -0.042266 | 2.2384
[13 empty] | cubic | linear
[12 empty] | Hi-tec | Carrier 2
[12 empty] | Calibration | Low
[12 empty] | names | Air
[12 empty] | A | -0.22611 | 0.035986
[12 empty] | Bx | 11.5842 | 10.6641
[12 empty] | Cx^2 | -0.56371
[12 empty] | Dx^3 | 0.084573
[12 empty] | Calib Temp Deg C | 17.4347
[12 empty] | Calib Press Mbars | 999.4546
[12 empty] | Date | 2015-02-02
Using domestic gas meter with a 1m long 22mm cooper pipe upstream to avoid jetting into meter